## Checklist
Centrifuge Operation SOP checklist
SOP-LAB-011 ¦ Rev 1.0 ¦ Owner: Laboratory Supervisor
Completed by: |  |  | Date: |  |  | Shift:
Step | Task | Details | Owner | Done | Initials | Time | Notes
5.1 | Inspect the centrifuge before use | Check the chamber, rotor and lid gasket for cracks, corrosion or debris, and confirm the unit is level and unobstructed before starting. | Lab Technician
5.2 | Select the correct rotor | Confirm the rotor installed matches the tube size and the speed and RCF required by the test method. | Lab Technician
5.3 | Weigh and balance tubes | Weigh each filled tube on the balance and pair it with a tube of matching weight directly opposite it in the rotor, using a balance blank if an odd number of samples is run. Checkpoint: Opposing tube pairs are within the rotor manufacturer's stated weight tolerance. | Lab Technician
5.4 | Load tubes symmetrically | Place balanced tube pairs into opposite rotor positions, filling the rotor symmetrically rather than clustering tubes on one side. Warning: Never run the centrifuge with an unbalanced or partially loaded rotor pattern. | Lab Technician
5.5 | Secure the rotor lid and chamber lid | Close and lock the rotor lid, then close and latch the centrifuge chamber lid before entering any run parameters. | Lab Technician
5.6 | Set run parameters | Enter the speed or RCF, run time, and temperature specified by the test method, confirming the settings match the protocol before starting. | Lab Technician
5.7 | Start the run and monitor briefly | Start the centrifuge and stay near it for the first minute of acceleration to listen and watch for vibration, unusual noise or an error code. Checkpoint: No unusual vibration or noise is detected during the initial acceleration. | Lab Technician
5.8 | Stop the run if imbalance occurs | If vibration, noise or an imbalance error occurs, stop the run immediately, wait for a complete stop, and re-balance the load before restarting. | Lab Technician
5.9 | Wait for a complete stop before opening | Wait until the rotor has come to a full stop, confirmed by the display or by listening, before unlatching the chamber lid. Warning: Never attempt to slow or stop a spinning rotor by hand or by opening the lid early. | Lab Technician
5.10 | Remove tubes carefully | Open the rotor lid, remove tubes without disturbing the separated layers, and place them upright in a rack for the next processing step. | Lab Technician
5.11 | Clean the rotor and chamber | Wipe down the rotor, buckets and chamber interior after each use, especially after any spill or tube breakage, per the manufacturer's cleaning instructions. | Lab Technician
5.12 | Log usage and report issues | Record the run parameters and operator in the usage log, and report any vibration, noise or damage to maintenance for follow-up. | Lab Technician
Supervisor sign-off: |  |  | Date:

## SOP
Centrifuge Operation SOP
Standard operating procedure
Document no. | SOP-LAB-011
Revision | 1.0
Owner | Laboratory Supervisor
Effective date | September 15, 2026
Review cycle | Every 12 months
1. Purpose
To operate laboratory centrifuges safely and consistently so that samples are processed correctly and equipment damage or injury from imbalance is prevented.
2. Scope
Applies to benchtop and floor-model centrifuges used to separate samples by density. Does not cover ultracentrifuges, which require manufacturer-specific training.
Definitions
Rotor | The removable component that holds sample tubes and spins inside the centrifuge.
Imbalance | An uneven weight distribution across the rotor that can cause vibration, damage or a safety hazard.
RCF | Relative centrifugal force, the effective spin force applied to a sample, often used instead of RPM in test methods.
3. Responsibilities
Lab Technician | Loads, balances and operates the centrifuge, and reports vibration or unusual noise.
Lab Supervisor | Approves protocol parameters and investigates reported imbalance or damage events.
Maintenance Technician | Services the centrifuge, replaces worn rotors or seals, and clears it to return to use.
RACI matrix
Inspect centrifuge before use | Lab Technician: R/A ¦ Lab Supervisor: I ¦ Maintenance Technician: -
Balance and load sample tubes | Lab Technician: R/A ¦ Lab Supervisor: - ¦ Maintenance Technician: -
Set and approve run parameters | Lab Technician: R ¦ Lab Supervisor: A ¦ Maintenance Technician: -
Service or repair the centrifuge | Lab Technician: I ¦ Lab Supervisor: C ¦ Maintenance Technician: R/A
4. Materials and PPE
Materials, tools and systems
Analytical or top-loading balance for tube balancing
Matched sample tubes and rotor buckets
Rotor and chamber cleaning supplies
Centrifuge usage and maintenance log
Manufacturer's operating manual for each rotor
Personal protective equipment
Disposable nitrile gloves
Safety glasses
Lab coat
5. Procedure
5.1 Inspect the centrifuge before use | Check the chamber, rotor and lid gasket for cracks, corrosion or debris, and confirm the unit is level and unobstructed before starting. Owner: Lab Technician
5.2 Select the correct rotor | Confirm the rotor installed matches the tube size and the speed and RCF required by the test method. Owner: Lab Technician
5.3 Weigh and balance tubes | Weigh each filled tube on the balance and pair it with a tube of matching weight directly opposite it in the rotor, using a balance blank if an odd number of samples is run. Owner: Lab Technician Checkpoint: Opposing tube pairs are within the rotor manufacturer's stated weight tolerance.
5.4 Load tubes symmetrically | Place balanced tube pairs into opposite rotor positions, filling the rotor symmetrically rather than clustering tubes on one side. Owner: Lab Technician Warning: Never run the centrifuge with an unbalanced or partially loaded rotor pattern.
5.5 Secure the rotor lid and chamber lid | Close and lock the rotor lid, then close and latch the centrifuge chamber lid before entering any run parameters. Owner: Lab Technician
5.6 Set run parameters | Enter the speed or RCF, run time, and temperature specified by the test method, confirming the settings match the protocol before starting. Owner: Lab Technician
5.7 Start the run and monitor briefly | Start the centrifuge and stay near it for the first minute of acceleration to listen and watch for vibration, unusual noise or an error code. Owner: Lab Technician Checkpoint: No unusual vibration or noise is detected during the initial acceleration.
5.8 Stop the run if imbalance occurs | If vibration, noise or an imbalance error occurs, stop the run immediately, wait for a complete stop, and re-balance the load before restarting. Owner: Lab Technician
5.9 Wait for a complete stop before opening | Wait until the rotor has come to a full stop, confirmed by the display or by listening, before unlatching the chamber lid. Owner: Lab Technician Warning: Never attempt to slow or stop a spinning rotor by hand or by opening the lid early.
5.10 Remove tubes carefully | Open the rotor lid, remove tubes without disturbing the separated layers, and place them upright in a rack for the next processing step. Owner: Lab Technician
5.11 Clean the rotor and chamber | Wipe down the rotor, buckets and chamber interior after each use, especially after any spill or tube breakage, per the manufacturer's cleaning instructions. Owner: Lab Technician
5.12 Log usage and report issues | Record the run parameters and operator in the usage log, and report any vibration, noise or damage to maintenance for follow-up. Owner: Lab Technician
6. Quality checks
Every load is balanced within the rotor's stated weight tolerance before a run starts.
The rotor lid and chamber lid are confirmed locked before each run.
No run is left unattended during the initial acceleration phase.
Reported vibration or noise events are logged and reviewed by maintenance.
7. Records
Centrifuge usage log
Tube balancing weight records
Maintenance and service history
Imbalance or incident reports
8. KPIs
Number of imbalance stops per month
Percentage of runs completed without interruption
Average centrifuge downtime for repair
Number of tube breakages per quarter
9. Common mistakes
Estimating tube balance by eye instead of weighing on a balance.
Opening the lid before the rotor has fully stopped.
Running a partial load without a balance blank in the opposite position.
Not reporting minor vibration because the run finished anyway.
10. Revision history
1.0 | Date: September 15, 2026 ¦ Description: Initial release ¦ Reviewed by: Ilia Pirozhenko

## Training Log
Centrifuge Operation SOP training log
Record every employee trained on this SOP.
Employee name | Role | Date trained | Trainer | Employee signature | Trainer signature | Notes

## Revision Log
Centrifuge Operation SOP revision log
Revision | Date | Description | Changed by | Approved by
1.0 | September 15, 2026 | Initial release | Ilia Pirozhenko

## About
About this template
This checklist comes from the free SOP template library by Perfect Wiki.
Open the template page
Keep this SOP up to date where your team works
Add it to Perfect Wiki so your team can find it and ask questions about it in Microsoft Teams, Slack, kChat or Mattermost.
Add to Perfect Wiki
This is a template. Adapt it to your organization, equipment and local regulations before use.