## Checklist
Pipette Calibration SOP checklist
SOP-LAB-006 ¦ Rev 1.0 ¦ Owner: QA Manager
Completed by: |  |  | Date: |  |  | Shift:
Step | Task | Details | Owner | Done | Initials | Time | Notes
5.1 | Check the calibration due date | Check the label on the pipette or the CMMS asset register to confirm it is due, and remove it from the bench rotation for testing. | Lab Technician
5.2 | Record room conditions | Record the ambient temperature, humidity and barometric pressure in the calibration log, since these affect the water density calculation. | Calibration Technician
5.3 | Inspect the pipette for damage | Check the shaft, tip cone, piston and seals for cracks, debris or wear, and clean the exterior before testing. | Calibration Technician
5.4 | Perform a leak test | Attach a tip, draw up liquid, and hold the pipette vertically for 20 to 30 seconds to check for a falling meniscus, which indicates a seal leak. Warning: Do not proceed to gravimetric testing on a pipette that fails the leak test; tag it out of service. | Calibration Technician
5.5 | Set the balance and tare | Place the weighing vessel on the analytical balance, close the draft shield, and tare the balance to zero before each dispense. | Calibration Technician
5.6 | Dispense at minimum, mid and maximum volume | Dispense the target volume into the weighing vessel and record the weight for at least ten replicates at each of the minimum, mid and maximum volumes for that pipette. Checkpoint: At least ten replicate weighings are recorded at each test volume before calculations begin. | Calibration Technician
5.7 | Calculate accuracy and precision | Convert each weight to volume using the water density at the recorded temperature, then calculate the mean volume, percent accuracy and percent CV for each test volume. | Calibration Technician
5.8 | Compare results to tolerance | Compare the calculated accuracy and precision against the manufacturer's stated tolerance for that pipette model and volume. | Calibration Technician
5.9 | Adjust or tag out of service if failing | If results are out of tolerance, adjust the pipette per the manufacturer's instructions and retest, or tag it out of service and arrange vendor service. | Calibration Technician
5.10 | Apply the pass or fail label | Attach a label to the pipette showing the calibration date, due date, result and technician initials. | Calibration Technician
5.11 | Log results and notify the owner | Enter the results into the calibration log or CMMS and notify the pipette's regular user that it is back in service or remains out for repair. | Calibration Technician
5.12 | Review failure trends | Review pipettes with repeated failures or adjustments across calibration cycles and decide whether replacement is more cost-effective than continued repair. | QA Manager
Supervisor sign-off: |  |  | Date:

## SOP
Pipette Calibration SOP
Standard operating procedure
Document no. | SOP-LAB-006
Revision | 1.0
Owner | QA Manager
Effective date | September 15, 2026
Review cycle | Every 12 months
1. Purpose
To confirm each pipette delivers volumes within its stated accuracy and precision so that dilutions, dosing and test results made with it stay reliable.
2. Scope
Applies to all single-channel and multichannel air-displacement pipettes used for testing or sample preparation. General instrument calibration is covered by the equipment calibration SOP.
Definitions
Gravimetric test | A calibration method that weighs dispensed water on an analytical balance to calculate the actual volume delivered.
Accuracy | How close the average delivered volume is to the target volume.
Precision (CV%) | How consistent repeated deliveries are with each other, expressed as a coefficient of variation.
3. Responsibilities
Lab Technician | Submits pipettes for calibration on schedule and reports suspected inaccuracy between calibrations.
Calibration Technician | Performs the gravimetric test, records results and applies the pass or fail label.
QA Manager | Sets tolerance criteria, reviews failed results and approves pipettes returning to service.
RACI matrix
Submit pipette for scheduled calibration | Lab Technician: R/A ¦ Calibration Technician: I ¦ QA Manager: -
Perform gravimetric calibration test | Lab Technician: I ¦ Calibration Technician: R/A ¦ QA Manager: C
Approve pipette adjustment or service | Lab Technician: I ¦ Calibration Technician: R ¦ QA Manager: A
Review calibration records and trends | Lab Technician: - ¦ Calibration Technician: R ¦ QA Manager: A
4. Materials and PPE
Materials, tools and systems
Analytical balance with calibration weights
Distilled or deionized water and thermometer
Calibration weights table for water density at temperature
Manufacturer-approved pipette tips
Calibration log or CMMS asset register
Pass or fail calibration labels
Personal protective equipment
Disposable nitrile gloves
5. Procedure
5.1 Check the calibration due date | Check the label on the pipette or the CMMS asset register to confirm it is due, and remove it from the bench rotation for testing. Owner: Lab Technician
5.2 Record room conditions | Record the ambient temperature, humidity and barometric pressure in the calibration log, since these affect the water density calculation. Owner: Calibration Technician
5.3 Inspect the pipette for damage | Check the shaft, tip cone, piston and seals for cracks, debris or wear, and clean the exterior before testing. Owner: Calibration Technician
5.4 Perform a leak test | Attach a tip, draw up liquid, and hold the pipette vertically for 20 to 30 seconds to check for a falling meniscus, which indicates a seal leak. Owner: Calibration Technician Warning: Do not proceed to gravimetric testing on a pipette that fails the leak test; tag it out of service.
5.5 Set the balance and tare | Place the weighing vessel on the analytical balance, close the draft shield, and tare the balance to zero before each dispense. Owner: Calibration Technician
5.6 Dispense at minimum, mid and maximum volume | Dispense the target volume into the weighing vessel and record the weight for at least ten replicates at each of the minimum, mid and maximum volumes for that pipette. Owner: Calibration Technician Checkpoint: At least ten replicate weighings are recorded at each test volume before calculations begin.
5.7 Calculate accuracy and precision | Convert each weight to volume using the water density at the recorded temperature, then calculate the mean volume, percent accuracy and percent CV for each test volume. Owner: Calibration Technician
5.8 Compare results to tolerance | Compare the calculated accuracy and precision against the manufacturer's stated tolerance for that pipette model and volume. Owner: Calibration Technician
5.9 Adjust or tag out of service if failing | If results are out of tolerance, adjust the pipette per the manufacturer's instructions and retest, or tag it out of service and arrange vendor service. Owner: Calibration Technician
5.10 Apply the pass or fail label | Attach a label to the pipette showing the calibration date, due date, result and technician initials. Owner: Calibration Technician
5.11 Log results and notify the owner | Enter the results into the calibration log or CMMS and notify the pipette's regular user that it is back in service or remains out for repair. Owner: Calibration Technician
5.12 Review failure trends | Review pipettes with repeated failures or adjustments across calibration cycles and decide whether replacement is more cost-effective than continued repair. Owner: QA Manager
6. Quality checks
Every calibration record shows at least ten replicates per tested volume.
Room temperature is recorded and used in the volume calculation.
Pipettes that fail the leak test are not gravimetrically tested until repaired.
Pass or fail labels match the result recorded in the calibration log.
7. Records
Calibration log or CMMS asset register
Gravimetric test data sheets
Pass or fail calibration labels
Vendor service records
8. KPIs
Percentage of pipettes calibrated on schedule
First-time pass rate on gravimetric testing
Number of pipettes requiring adjustment per cycle
Average pipette downtime for repair
9. Common mistakes
Skipping the leak test and going straight to gravimetric testing.
Not recording room temperature before calculating volumes.
Testing with tips that are not the manufacturer-approved brand for that pipette.
Using fewer than the required number of replicates per volume.
10. Revision history
1.0 | Date: September 15, 2026 ¦ Description: Initial release ¦ Reviewed by: Ilia Pirozhenko

## Training Log
Pipette Calibration SOP training log
Record every employee trained on this SOP.
Employee name | Role | Date trained | Trainer | Employee signature | Trainer signature | Notes

## Revision Log
Pipette Calibration SOP revision log
Revision | Date | Description | Changed by | Approved by
1.0 | September 15, 2026 | Initial release | Ilia Pirozhenko

## About
About this template
This checklist comes from the free SOP template library by Perfect Wiki.
Open the template page
Keep this SOP up to date where your team works
Add it to Perfect Wiki so your team can find it and ask questions about it in Microsoft Teams, Slack, kChat or Mattermost.
Add to Perfect Wiki
This is a template. Adapt it to your organization, equipment and local regulations before use.